Portable Multi-   parameter Water

Quality Analyzer

User Manual

# Contents

Introduction 	3

1 Overview 	4

1.1 Technical parameters 	4

2 Operations 	6

2.1 Key Description  	6

2.2 Measurement Interface 	6

2.3  Setup  Screen 	7

2.4 Calibration Interface 	7

3 Maintenance 	11

3.1 Meter Maintenance 	11

4 Sensor Maintenance 	11

5 FAQs 	14

# Introduction

Dear user

Thank you  so  much  for purchasing  the portable multi-parameter water  quality analyzer. Before you use, please read this manual in detail, it will be of great help to the use and maintenance of this instrument, and can avoid unnecessary trouble due to

improper operation and maintenance.

Please abide by the operating procedures and precautions in this manual.

In order to ensure the effective after-sales protection provided by this instrument, please do not use and maintain this instrument in ways other than those specified in this

manual.

Any failure or loss caused by failure to comply with the precautions specified in this manual is not covered by the manufacturer's warranty, and the manufacturer does not assume any related responsibility. Please take good care of all documents. If you

have any questions, please contact our after-sales service department.

Upon  receipt  of the  instrument,  please  open  the  package  carefully  and  check whether the instrument and accessories are damaged during delivery. If any damage is found, please contact our after-sales service department and keep the packaging for

return.

When the instrument malfunctions, please do not repair it yourself. Please contact

our after-sales service department.

# 1 Overview

Portable multi-parameter water quality analyzer measure temperature, pH, COD, TSS and NH3-N.  These  data  can be  automatically recognized when  connected to  a transmitter.  The  instrument  has  the  operating  keyboard  key.  The   function   of  the instrument is comprehensive. The operation and the interface are simple. Moreover, the

instrument can realize the sensor calibration function.

## 1 . 1 Technical parameters

| Handheld Multi-parameter Probe Specifications |  |  |
| --- | --- | --- |
| pH Sensor | Range | 0 to 14 |
|  | Precision | 0.02 |
|  | Resolution | 0.01 |
|  | Calibrate | Two-point calibration |
| NH3-N sensor | Range | 0- 100mg/L |
|  | Precision | ±10% or ±1mg/L |
|  | Resolution | 0.01 mg/L |
|  | Calibrate | Two-point calibration |
| COD sensor | Measuring Range | 0- 1500mg/L |
|  | Precision | ±10% |
|  | Resolution | 0. 1 mg/L |
|  | Calibrate | Two-point calibration |
| TSS sensor | Range | 0-4000mg/L |
|  | Resolution | 0. 1 |
|  | Precision | ±5% |
|  | Calibration | One-point calibration |

| Temperature | Range | 0~50 ° C (no freezing) |
| --- | --- | --- |
|  | Precision | ±0 .5 ℃ |
|  | Resolution | 0. 1 ℃ |
| Sleep Timer | Support |  |

Short press: less than 2s Long press: more than 3s

2 Operations

Long  press  the   "OK"  key  to  power   on.  When  no  probe   is  connected,  the measurement  interface  displays   0.  When  the  probe  is  inserted,  the  measurement

interface displays all the data of the probe.

## 2. 1 Key Description

| 1 |  | Long press | On/off |
| --- | --- | --- | --- |
| 2 |  | Short press | Access the menu screen |
| 3 |  | Short press | Exit Settings Return to the previous level |
| 4 |  | Short press |  |
| 5 |  | Short press | Scroll down to check the menu options Reduce the value when you set |

## 2 .2 Measurement Interface

NH3-N  043.5mg/L

Temp 29. 1℃

pH  06.7

COD  0043.5mg/l

TSS   0029.0mg/l

The  measurement  interface  in the    left    picture    contains    4 parameters  of  probes:  NH3-N, COD, TSS, pH and Temp.Press the  Down  button  to   flip   the

screen.

2.3 Setup Screen

Main

1 System Settings

2 Signal Settings

3 On-line Calibration

## 2 .4 Calibration Interface

Press   "MENU "   to   enter   the   main

menu .

The  system  setting  interface can  set the  language  type.  Signal   Settings are  used  to  specify  the  parameter Settings of the sensor .

On-line     calibration     is     used     to

calibrate the sensor .

Select a probe under Menu-> Main Menu-> Online Calibration directory to enter

the calibration interface. Do not unplug the sensor cable during the calibration process.

(1) TSS calibration

TSS sensor supports two-point calibration.

Configure 2000mg/L of standard liquid, place the sensor in it, and click "Zero" to calibrate after stabilizing. Configure 4000mg/L of the standard liquid and place the

sensor in it. After stabilizing, click "Slope" to calibrate.

(2) pH calibration

pH electrodes support two-point calibration. And the two calibration point liquids

should be fixed 4.01, 7.00.

Put the sensor into the calibration liquid with pH=7. After the reading is stable, select  "midpoint"  and  click  OK  to  complete  the  first  calibration  point.  Rinse  the electrode head with clean water and put the sensor into the calibration solution with pH=4.01. After the reading is stable, select "low point" and click OK to complete the

second point calibration.

(3) Calibration of ammonia nitrogen sensor

Ammonia nitrogen supports two-point calibration.

Configure  10mg/L of standard liquid, place the sensor in it, and click "zero" to calibrate  after  stabilizing.  Configure  100mg/L  of the  standard  liquid  and  place  the

sensor in it. After stabilizing, click "Slope" to calibrate.

(4) COD sensor calibration

COD supports two-point calibration.

Configure  50mg/L  standard  liquid,  place  the  sensor  in  it,  and  click  "zero"  to calibrate  after  stabilizing.  Configure  800mg/L  of the  standard  liquid  and  place  the

sensor in it. After stabilizing, click "Slope" to calibrate.

# 3  Maintenance

## 3 . 1 Meter Maintenance

No  other  maintenance  is  required  except  the  occasional  wiping  of  the  outer surface of the meter with a damp cloth. Usually use the meter should be careful not to

water.

# 3 .2 Sensor Maintenance

TSS Sensor

In  order  to  ensure  the  accuracy  of the  measurement,  the  sensor  needs  to  be regularly  maintained  and  maintained.  Including  the  cleaning  of  the  sensor,  the

appearance of the sensor is damaged, etc.

1) Sensor cleaning and inspection

Due  to  more  impurities  in  the  water  quality  environment,  the  probe  may be polluted by algae after a period of use, and it needs to be cleaned regularly. There is no need to remove the sensor during the cleaning process, use a wet and soft cloth or water sponge to gently wipe the measuring window, and for some stubborn dirt, you can add some household detergent to the water for cleaning. Note: The maintenance frequency  is  recommended  to  arrange  the  cleaning  cycle  according  to  the  water condition. Check whether the appearance of the sensor is damaged. If the shell or cable

is damaged, stop using the sensor and contact the after-sales service for replacement.

2) Storage of the sensor

When not in use, the sensor is powered off and dried for preservation.

NH3-N sensor

| Maintenance tasks | Suggested frequency of maintenance |
| --- | --- |
| Cleaning the sensor | Clean regularly |
| Calibrate the sensor | Over the measurement error , perform the two-point calibration |

Maintenance methods:

1）Clean the outer surface of the sensor: clean the outer surface of the sensor with tap water, if there is still debris remaining, wipe with a wet soft cloth, for some

stubborn dirt, you can add some household washing liquid in tap water to clean.

Check whether the terminal is dry. Otherwise, please wipe with anhydrous alcohol,

blow dry after use.

2）Check the cable of the sensor: if the shell or cable is damaged, it should stop using and contact the after-sales service for replacement in time. The cable should not be taut during normal work, otherwise it is easy to break the internal wire of the cable,

causing the sensor to not work normally.

3）Sensor storage: The electrodes are not used for a long time (more than two weeks) should be stored dry. The electrodes induction elements should be covered in the

protective cap.

pH electrode

| Maintenance tasks | Suggested frequency of maintenance |
| --- | --- |
| Calibrate the sensor | Perform the two-point calibration based on the measurement |
| Change the pH electrode | Please replace the electrode according to the warranty period and application environment |

Maintenance method:

1)  Clean the outer surface of the sensor: Observe the glass measuring part of the electrode head. If the glass contact surface is covered, use a wet cotton swab and

gently wipe the surface until it is clean!

2)  Check the cable of the sensor: If the shell or cable is damaged, stop using it and

contact the after-sales service for replacement in time.

3)  Sensor storage: When the sensor is not in use, the black plastic cap should be fastened to check whether the sponge inside is wet, if it is not wet enough, please add 3mol/L of potassium chloride solution, and let the electrode be stored in the

solution with potassium chloride.

COD sensor

| Maintenance tasks | Suggested frequency of maintenance |
| --- | --- |
| Cleaning the sensor | Clean regularly |

| Calibrate the sensor | Over the measurement error, perform the two- point calibration |
| --- | --- |

1）Clean the outer surface of the sensor: Clean the outer surface of the sensor with tap water, if there is still debris residue, wipe with a wet soft cloth, for some stubborn dirt, you can add some household washing liquid in tap water to clean. Check whether the window of the sensor is dirt and whether the cleaning brush is

working properly.

2）Check the cable of the sensor: If the shell or cable is damaged, it should stop using and contact the after-sales service for replacement in time. The cable should not be tense during normal work, otherwise it is easy to break the internal wire of

the cable, causing the sensor to not work normally.

3）Sensor storage: The electrodes are not used for a long time (more than two weeks)

should be stored dry.

Precautions:

All sensors contain sensitive optical and electronic components. Make sure the probe is protected from violent mechanical shocks. Damage to the sensor due to impact

or human caused will not be covered by the warranty.

# 4 FAQs

Table 5- 1 lists the possible issues and solutions for handheld instruments. Please

contact with us if your problem is not listed or the solution cannot handle your problem.

Table 5- 1 Lists of frequently asked questions.

| Abnormal phenomenon | Possible causes | Solutions |
| --- | --- | --- |
| Handheld instrument does not display measurement results, when the sensor is connected. | The meter is incorrectly connected to the cable . | Reconnect the meter and the cable . |
|  | The meter battery is flat. | Charge the battery after shutdown . |
|  | Cable failure | Please contact us . |
| The measured value is too high , too low , or the value is continuously unstable. | The measuring end face of probe whether have bubbles or dirt attachment. | Check and remove . |
|  | The consumables such as pH electrodes are abnormal. | Check and replace . |
|  | The probes are not calibrated on schedule . | Perform calibration operations. |
|  | The data of the other probes is abnormal. | Please contact us . |